## REFERENCE SHEET
Electricity Cost/kWh (£) | 0.1
T/CO2/kWh | 0.00028307
Cost T/CO2 (£) | 18

## Replacement units
Unit | Size | Number Bought | kWh/hr | kWh/day | kWh/yr | kWh/Yr (£) | T/CO2/Yr | T/CO2 (£) | Hr/Day | Days/Year | Total kWh/yr | Total kWh/Yr (£) | Total T/CO2/Yr | Total T/CO2 (£)
E3 100 | 100 | 39 | 0.232 | 2.784 | 668.16 | 66.816 | 0.18914 | 3.4044 | 12 | 240 | 26058.24 | 2605.824 | 7.3763 | 132.7735
E3 190 | 190 | 75 | 0.3467 | 4.1604 | 998.496 | 99.8496 | 0.28264 | 5.0876 | 12 | 240 | 74887.2 | 7488.72 | 21.1983 | 381.5698
E3 200 | 200 | 13 | 0.36 | 4.32 | 1036.8 | 103.68 | 0.29349 | 5.2828 | 12 | 240 | 13478.4 | 1347.84 | 3.8153 | 68.676
E3 200/F | 200 | 1 | 0.54 | 6.48 | 1555.2 | 155.52 | 0.44023 | 7.9241 | 12 | 240 | 1555.2 | 155.52 | 0.44023 | 7.9241
E3 425 | 425 | 4 | 0.553 | 6.636 | 1592.64 | 159.264 | 0.45083 | 8.1149 | 12 | 240 | 6370.56 | 637.056 | 1.8033 | 32.4597
E3 425/F | 425 | 1 | 0.763 | 9.156 | 2197.44 | 219.744 | 0.62203 | 11.1965 | 12 | 240 | 2197.44 | 219.744 | 0.62203 | 11.1965
E3 885/F | 885 | 21 | 1.132 | 13.584 | 3260.16 | 326.016 | 0.92285 | 16.6114 | 12 | 240 | 68463.36 | 6846.336 | 19.3799 | 348.8386
[11 empty] | 193010.4 | 19301.04 | 54.6355 | 983.4382
Cost of Project | 210000
RHS 190 | 41
LHS 190 | 34

## Exisiting Units
Unit | Manufacturer | Model | Location | Department | Net Capacity (L) | kWh/Hr | (kWh/day) | (kWh/year) | Electricity (£) 1 Yr | CO2 tonnes/yr | CO2 cost (£) | Total Cost/yr (£) | Set Temperature (C) | Items stored on top | Hrs/Day | Days/Yr
DC 147 | LEEC | Fan | Begbroke | Begbroke | 1000 | 3.12 | 74.88 | 27331.2 | 2733.12 | 7.7366 | 139.2596 | 2949.2595 | 75 |  | 24 | 365
DC 1 | LTE unitemp | Fan | 00-052 | Biochemistry | 1000 | 2.7267 | 65.44 | 23885.6 | 2388.56 | 6.7613 | 121.7033 | 2577.4511 | 65 |  | 24 | 365
DC 2 | LTE unitemp | Fan | 00-052 | Biochemistry | 1000 | 2.7267 | 65.44 | 23885.6 | 2388.56 | 6.7613 | 121.7033 | 2577.4511 | 65 |  | 24 | 365
DC 3 | LTE unitemp | Fan | 00-029 | Biochemistry | 534 | 1.686 | 40.4651 | 14769.7674 | 1476.9767 | 4.1809 | 75.2558 | 1593.7784 | 60 |  | 24 | 365
DC 4 | Genlab | IWC425 Convection | 20-059 | Biochemistry | 425 | 0.33306 | 7.9935 | 2917.6163 | 291.7616 | 0.82589 | 14.866 | 314.8346 | 45 |  | 24 | 365
DC 5 | LTE | Convection | 30-008 | Biochemistry | 100 | 0.63544 | 15.2506 | 5566.481 | 556.6481 | 1.5757 | 28.3627 | 600.6687 | 75 |  | 24 | 365
DC 6 | LEEC | Convection | 40-039 | Biochemistry | 100 | 0.33799 | 8.1118 | 2960.8035 | 296.0803 | 0.83811 | 15.0861 | 319.4949 | 45 |  | 24 | 365
DC 7 | LTE | Convection | 40-039 | Biochemistry | 200 | 0.33799 | 8.1118 | 2960.8035 | 296.0803 | 0.83811 | 15.0861 | 319.4949 | 75 |  | 24 | 365
DC 8 | LTE | Convection | 20.16 | Biochemistry | 200 | 1.0086 | 24.2057 | 8835.0857 | 883.5086 | 2.5009 | 45.0171 | 953.3778 | 75 |  | 24 | 365
DC 141 | Genlab | Fan | F29 | Botnar Building | 1000 | 1.9215 | 46.1153 | 16832.0982 | 1683.2098 | 4.7647 | 85.7639 | 1816.3207 | 64 |  | 24 | 365
DC 146 | LTE Harvard | Convection | CBF | CBF | 450 | 0.39667 | 3.1733 | 1158.2667 | 115.8267 | 0.32787 | 5.9017 | 375 | 55 |  | 8 | 365
DC 145 | Genlab | Fan DC1000 | 100 F13 | CCVTM | 1000 | 0.91 | 21.84 | 7971.6 | 797.16 | 2.2565 | 40.6174 | 860.2007 | 40 |  | 24 | 365
DC 135 | LEEC | Convection | 152.30.53 | Centre for the Environment | 100 | 0.17818 | 4.2764 | 1560.8727 | 156.0873 | 0.44184 | 7.9531 | 168.4309 | 40 | X | 24 | 365
DC 136 | LEEC | Convection | 152.10.93a | Centre for the Environment | 200 | 0.74819 | 17.9566 | 6554.1687 | 655.4169 | 1.8553 | 33.3952 | 707.2483 | 60 |  | 24 | 365
DC 137 | LEEC | Convection | 152.10.34 | Centre for the Environment | 100 | 0.501 | 12.024 | 4388.76 | 438.876 | 1.2423 | 22.3619 | 473.583 | 75 |  | 24 | 365
DC 138 | Victor | Convection | 152.10.33 | Centre for the Environment | 100 | 0.501 | 12.024 | 4388.76 | 438.876 | 1.2423 | 22.3619 | 473.583 | 75 |  | 24 | 365
DC 139 | LEEC | Convection | 152.10.35 | Centre for the Environment | 100 | 0.51097 | 12.2632 | 4476.0774 | 447.6077 | 1.267 | 22.8068 | 483.0053 | 70 |  | 24 | 365
DC 100 | LEEC | Convection | G Floor Lab 2 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 101 | LEEC | Convection | G Floor Lab 3 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 102 | LEEC | Convection | G Floor Lab 4 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 103 | LEEC | Convection | G Floor Lab 4 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 104 | LEEC | Convection | G Floor Lab 5 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 105 | LEEC | Convection | G Floor Lab 5 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 106 | LEEC | Convection | G Floor Lab 6 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 107 | LEEC | Convection | G Floor Lab 6 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 108 | LEEC | Convection | LG Floor Lab 8 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 109 | LEEC | Convection | LG Floor Lab 8 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 110 | LEEC | Convection | LG Floor Lab 9 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 111 | LEEC | Convection | LG Floor Lab 9 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 112 | LEEC | Convection | LG Floor Lab 10 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 113 | LEEC | Convection | LG Floor Lab 10 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 114 | LEEC | Convection | LG Floor Lab 11 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 115 | LEEC | Convection | LG Floor Lab 11 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 |  | 17 | 365
DC 116 | LEEC | Convection | LG Floor Lab 1 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 117 | LEEC | Convection | LG Floor Lab 2 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 |  | 17 | 365
DC 118 | LEEC | Convection | B Darkroom 00.006 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 |  | 17 | 365
DC 119 | LEEC | Convection | B 5&6 Prep. Area | Chemistry Research Lab | 1000 | 2.3825 | 57.18 | 20870.7 | 2087.07 | 5.9079 | 106.3416 | 2252.1188 | 65 |  | 24 | 365
DC 38 | LEEC | Convection | 2nd Floor Lab 6 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 |  | 17 | 365
DC 39 | LEEC | Convection | 2nd Floor Lab 5 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 |  | 17 | 365
DC 40 | LEEC | Convection | 2nd Floor Lab 5 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 41 | LEEC | Convection | 2nd Floor Lab 4 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 42 | LEEC | Convection | 2nd Floor Lab 4 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 |  | 17 | 365
DC 43 | LEEC | Convection | 2nd Floor Lab 3 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 |  | 17 | 365
DC 44 | LEEC | Convection | 2nd Floor Lab 3 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 |  | 17 | 365
DC 45 | LEEC | Convection | 2nd Floor Lab 2 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 46 | LEEC | Convection | 2nd Floor Lab 2 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 47 | LEEC | Convection | 2nd Floor Lab 1 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 |  | 17 | 365
DC 48 | LEEC | Convection | 2nd Floor Lab 1 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 49 | LEEC | Convection | 2nd Floor Lab 12 | Chemistry Research Lab | 100 | 0.48938 | 11.745 | 4286.925 | 428.6925 | 1.2135 | 21.843 | 462.5942 | 68 | X | 24 | 365
DC 50 | LEEC | Convection | 2nd Floor Lab 12 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 51 | LEEC | Convection | 2nd Floor Lab 12 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 52 | LEEC | Convection | 2nd Floor Lab 11 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 53 | LEEC | Convection | 2nd Floor Lab 11 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 |  | 17 | 365
DC 54 | LEEC | Convection | 2nd Floor Lab 10 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 |  | 17 | 365
DC 55 | LEEC | Convection | 2nd Floor Lab 10 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 |  | 17 | 365
DC 56 | LEEC | Convection | 2nd Floor Lab 9 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 57 | LEEC | Convection | 2nd Floor Lab 9 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 58 | LEEC | Convection | 2nd Floor Lab 8 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 59 | LEEC | Convection | 2nd Floor Lab 8 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 60 | LEEC | Convection | 2nd Floor Lab 7 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 |  | 17 | 365
DC 61 | LEEC | Convection | 2nd Floor Lab 7 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 |  | 17 | 365
DC 62 | LEEC | Convection | 1st Floor Lab 7 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 63 | LEEC | Convection | 1st Floor Lab 7 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 64 | LEEC | Convection | 1st Floor Lab 8 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 65 | LEEC | Convection | 1st Floor Lab 8 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 66 | LEEC | Convection | 1st Floor Lab 9 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 67 | LEEC | Convection | 1st Floor Lab 9 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 68 | LEEC | Convection | 1st Floor Lab 10 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 69 | LEEC | Convection | 1st Floor Lab 10 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 70 | LEEC | Convection | 1st Floor Lab 11 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 |  | 17 | 365
DC 71 | LEEC | Convection | 1st Floor Lab 11 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 |  | 17 | 365
DC 72 | LEEC | Convection | 1st Floor Lab 12 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 73 | LEEC | Convection | 1st Floor Lab 12 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 74 | LEEC | Convection | 1st Floor Lab 1 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 75 | LEEC | Convection | 1st Floor Lab 1 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 76 | LEEC | Convection | 1st Floor Lab 2 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 77 | LEEC | Convection | 1st Floor Lab 2 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 78 | LEEC | Convection | 1st Floor Lab 3 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 79 | LEEC | Convection | 1st Floor Lab 3 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 80 | LEEC | Convection | 1st Floor Lab 4 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 81 | LEEC | Convection | 1st Floor Lab 4 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 82 | LEEC | Convection | 1st Floor Lab 5 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 83 | LEEC | Convection | 1st Floor Lab 6 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 84 | LEEC | Convection | 1st Floor Lab 6 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 85 | LEEC | Convection | G Floor Lab 7 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 86 | LEEC | Convection | G Floor Lab 7 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 87 | LEEC | Convection | G Floor Lab 8 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 88 | LEEC | Convection | G Floor Lab 8 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 89 | LEEC | Convection | G Floor Lab 9 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 90 | LEEC | Convection | G Floor Lab 9 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 91 | LEEC | Convection | G Floor Lab 10 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 92 | LEEC | Convection | G Floor Lab 10 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 93 | LEEC | Convection | G Floor Lab 11 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 94 | LEEC | Convection | G Floor Lab 11 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 95 | LEEC | Convection | G Floor Lab 12 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 96 | LEEC | Convection | G Floor Lab 12 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 97 | LEEC | Convection | G Floor Lab 1 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 98 | LEEC | Convection | G Floor Lab 1 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 99 | LEEC | Convection | G Floor Lab 2 | Chemistry Research Lab | 191 | 0.45333 | 7.7067 | 2812.9333 | 281.2933 | 0.79626 | 14.3326 | 428.48 | 75 | X | 17 | 365
DC 140 | Belling | Convection | Chemistry Lab 1 | Dyson Perrins Building | 98 | 0.57 | 13.68 | 4993.2 | 499.32 | 1.4134 | 25.4417 | 538.807 | 75 | X | 24 | 365
DC 12 | LEEC | Convection | S13 | Inorganic Chemistry | 100 | 0.434 | 10.416 | 3801.84 | 380.184 | 1.0762 | 19.3714 | 410.2496 | 55 |  | 24 | 365
DC 13 | Belling | Convection | S13 | Inorganic Chemistry | 98 | 0.39479 | 9.4749 | 3458.3453 | 345.8345 | 0.97895 | 17.6212 | 373.1837 | 75-77 |  | 24 | 365
DC 14 | B&T | Convection | S7 | Inorganic Chemistry | 100 | 0.556 | 13.344 | 4870.56 | 487.056 | 1.3787 | 24.8168 | 525.5732 | 76 |  | 24 | 365
DC 15 | LEEC | Convection | S6 | Inorganic Chemistry | 100 | 0.502 | 12.048 | 4397.52 | 439.752 | 1.2448 | 22.4065 | 474.5283 | 58 |  | 24 | 365
DC 16 | Marose Scientific | Convection | S2 | Inorganic Chemistry | 113 | 0.67292 | 16.15 | 5894.75 | 589.475 | 1.6686 | 30.0353 | 636.0916 | 70 | X | 24 | 365
DC 17 | Gallenkamp Economy Incubator Size 2 | Convection | T1 | Inorganic Chemistry | 104 | 0.1843 | 4.4232 | 1614.471 | 161.4471 | 0.45701 | 8.2261 | 174.2146 | 50 |  | 24 | 365
DC 18 | Marose Scientific | Convection | T7 | Inorganic Chemistry | 113 | 0.75689 | 18.1654 | 6630.3604 | 663.036 | 1.8769 | 33.7834 | 715.47 | 75 |  | 24 | 365
DC 19 | LEEC | Convection | T16 | Inorganic Chemistry | 200 | 0.7278 | 17.4671 | 6375.509 | 637.5509 | 1.8047 | 32.4849 | 687.9694 | 60 | X | 24 | 365
DC 20 | LEEC | Convection | F22 | Inorganic Chemistry | 100 | 0.49615 | 11.9077 | 4346.3077 | 434.6308 | 1.2303 | 22.1456 | 469.0021 | 74 |  | 24 | 365
DC 21 | LEEC | Convection | F4a | Inorganic Chemistry | 100 | 0.47674 | 11.4419 | 4176.2791 | 417.6279 | 1.1822 | 21.2792 | 450.6546 | 59 | X | 24 | 365
DC 22 | LEEC | Convection | F floor Teaching Lab Section 1 | Inorganic Chemistry | 100 | 0.501 | 12.024 | 4388.76 | 438.876 | 1.2423 | 22.3619 | 473.583 | 75 |  | 24 | 365
DC 23 | LEEC | Convection | F floor Teaching Lab Section 2 | Inorganic Chemistry | 200 | 0.762 | 18.288 | 6675.12 | 667.512 | 1.8895 | 34.0115 | 720.2999 | 75 |  | 24 | 365
DC 24 | LTE | 322/0104/00 Convection | F floor Teaching Lab Section 2 | Inorganic Chemistry | 180 | 0.68082 | 16.3396 | 5963.951 | 596.3951 | 1.6882 | 30.3879 | 643.559 | 75 |  | 24 | 365
DC 25 | LEEC | Convection | F32 | Inorganic Chemistry | 200 | 0.74074 | 17.7778 | 6488.8889 | 648.8889 | 1.8368 | 33.0626 | 700.2041 | 68 |  | 24 | 365
DC 26 | Gallenkamp Hotbox Oven Size 1 | Convection | F13 | Inorganic Chemistry | 52 | 0.095 | 2.28 | 832.2 | 83.22 | 0.23557 | 4.2403 | 89.8012 | less than 65 |  | 24 | 365
DC 27 | LEEC | Convection | CG14 168.10.06 | Inorganic Chemistry | 100 | 0.501 | 12.024 | 4388.76 | 438.876 | 1.2423 | 22.3619 | 473.583 | 75 |  | 24 | 365
DC 28 | Corsair | Convection | G3 | Inorganic Chemistry | 100 | 0.44 | 10.56 | 3854.4 | 385.44 | 1.0911 | 19.6392 | 415.9212 | 60 |  | 24 | 365
DC 29 | LEEC | Convection | G2 | Inorganic Chemistry | 100 | 0.49867 | 11.968 | 4368.32 | 436.832 | 1.2365 | 22.2577 | 471.3774 | 75 |  | 24 | 365
DC 30 | LEEC | Convection | B5 | Inorganic Chemistry | 100 | 0.501 | 12.024 | 4388.76 | 438.876 | 1.2423 | 22.3619 | 473.583 | 75 |  | 24 | 365
DC 142 | Genlab | Fan | 668.10.24 | Kennedy Building | 1000 | 0.32667 | 2.6133 | 953.8667 | 95.3867 | 0.27001 | 4.8602 | 308.72 | 50 |  | 8 | 365
DC 143 | Genlab | Fan DC1000 | 664.10.18 | Nuffield Department of Medicine | 1000 | 1.1046 | 8.8367 | 2306.37 | 230.637 | 0.65286 | 11.7516 | 723.82 | 40 and 60 |  | 8 | 261
DC 144 | Genlab | Fan DC1000 | 664.10.18 | Nuffield Department of Medicine | 1000 | 1.1046 | 8.8367 | 2306.37 | 230.637 | 0.65286 | 11.7516 | 723.82 | 40 and 60 |  | 8 | 261
DC 124 | Genlab | Fan | 30.102 | Old Road Campus Research Building | 1000 | 1.09 | 26.16 | 9548.4 | 954.84 | 2.7029 | 48.6516 | 1030.3503 | 60 |  | 24 | 365
DC 125 | LEEC | Convection | 30.08 | Old Road Campus Research Building | 1000 | 2.3273 | 55.8545 | 20386.9091 | 2038.6909 | 5.7709 | 103.8766 | 2199.9138 | 68 |  | 24 | 365
DC 126 | Genlab | Fan | 20.107 | Old Road Campus Research Building | 1000 | 1.2145 | 29.1484 | 10639.1613 | 1063.9161 | 3.0116 | 54.2093 | 1148.0523 | 60 |  | 24 | 365
DC 127 | Genlab | Fan | 10.109 | Old Road Campus Research Building | 1000 | 1.15 | 27.6 | 10074 | 1007.4 | 2.8516 | 51.3296 | 1087.0668 | 60 |  | 24 | 365
DC 128 | Genlab | Fan | 00.105 | Old Road Campus Research Building | 1000 | 1.15 | 27.6 | 10074 | 1007.4 | 2.8516 | 51.3296 | 1087.0668 | 60 |  | 24 | 365
DC 10 | LEEC | Convection | 336.30.21 | Pathology | 100 | 0.29811 | 7.1547 | 2611.4717 | 261.1472 | 0.73923 | 13.3061 | 281.7991 | 50 | X | 24 | 365
DC 11 | LEEC | Convection | 336.30.21 | Pathology | 100 | 0.23019 | 5.5245 | 2016.4528 | 201.6453 | 0.5708 | 10.2744 | 217.5917 | 43 |  | 24 | 365
DC 9 | Leader | DC113 Convection | 214.10.25 | Pathology | 105 | 0.286 | 6.864 | 2505.36 | 250.536 | 0.70919 | 12.7655 | 270.3488 | 40 | X | 24 | 365
DC 31 | LEEC | Convection | 40.39 | Pharmacology | 200 | 0.635 | 15.24 | 5562.6 | 556.26 | 1.5746 | 28.3429 | 600.2499 | 75 |  | 24 | 365
DC 32 | Belling | Convection | 40.39 | Pharmacology | 98 | 0.54314 | 13.0353 | 4757.8824 | 475.7882 | 1.3468 | 24.2426 | 513.4143 | 87 |  | 24 | 365
DC 33 | Belling | Convection | 30.84 | Pharmacology | 98 | 0.56 | 13.44 | 4905.6 | 490.56 | 1.3886 | 24.9953 | 529.3543 | 75 |  | 24 | 365
DC 34 | LEEC | Convection | 30.18 | Pharmacology | 100 | 0.32781 | 7.8675 | 2871.6556 | 287.1656 | 0.81288 | 14.6318 | 309.8751 | 61 |  | 24 | 365
DC 35 | LEEC | Convection | 20.11 | Pharmacology | 200 | 0.30915 | 7.4197 | 2708.1763 | 270.8176 | 0.7666 | 13.7989 | 292.2343 | 41 |  | 24 | 365
DC 36 | Genlab | Convection | 10.5 | Pharmacology | 200 | 0.759 | 18.216 | 6648.84 | 664.884 | 1.8821 | 33.8776 | 717.4641 | 75 |  | 24 | 365
DC 37 | Belling | Convection | 10.07 | Pharmacology | 98 | 0.56 | 13.44 | 4905.6 | 490.56 | 1.3886 | 24.9953 | 529.3543 | 75 | X | 24 | 365
DC 120 | LEEC | Convection | S210 | Plant Sciences | 181 | 0.874 | 20.976 | 7656.24 | 765.624 | 2.1673 | 39.0105 | 826.1708 | 75 |  | 24 | 365
DC 121 | LTE | Convection | N305(A) | Plant Sciences | 200 | 0.89474 | 21.4737 | 7837.8947 | 783.7895 | 2.2187 | 39.9361 | 845.7728 | 54 |  | 24 | 365
DC 122 | LTE | 322/0104/00 Convection | N203 | Plant Sciences | 200 | 1.0158 | 24.3795 | 8898.5302 | 889.853 | 2.5189 | 45.3403 | 960.224 | 76 |  | 24 | 365
DC 123 | LEEC | Convection | N101 | Plant Sciences | 181 | 0.874 | 20.976 | 7656.24 | 765.624 | 2.1673 | 39.0105 | 826.1708 | 75 |  | 24 | 365
DC 131 | Genlab | Fan | 327 | Weatherall Institute of Molecular Medicine | 1000 | 0.96346 | 23.1231 | 8439.9231 | 843.9923 | 2.3891 | 43.0036 | 910.7366 | 50 | X | 24 | 365
DC 132 | LTE Unitemp | Convection | 450 | Weatherall Institute of Molecular Medicine | 1000 | 1.8189 | 43.6547 | 15933.9789 | 1593.3979 | 4.5104 | 81.1878 | 1719.4063 | 55 | X | 24 | 365
DC 133 | LEEC | F1 Convection | 403 | Weatherall Institute of Molecular Medicine | 1000 | 0.83736 | 20.0967 | 7335.2967 | 733.5297 | 2.0764 | 37.3752 | 791.5384 | 40 | X | 24 | 365
DC 134 | LEEC | F1 Convection | 403 | Weatherall Institute of Molecular Medicine | 1000 | 0.81099 | 19.4637 | 7104.2637 | 710.4264 | 2.011 | 36.1981 | 766.608 | 41 | X | 24 | 365
DC 129 | LTE Unitemp | Fan | 00/134 | Wellcome Trust Centre for Human Genetics | 1000 | 0.6825 | 4.095 | 1494.675 | 149.4675 | 0.4231 | 7.6158 | 447.21 | 65 | X | 6 | 365
DC 130 | LTE Unitemp | Convection | 10.3 | Wellcome Trust Centre for Human Genetics | 1000 | 1.5338 | 36.81 | 13435.65 | 1343.565 | 3.8032 | 68.4581 | 1449.8163 | 46 | X | 24 | 365
[8 empty] | 694119.5634 | 69411.9563 | 196.4844

## Project Summary
% Reduction in kWh | 72.1935
kWh/Yr Saved | 501109.1634
kWh/Yr Saved (£) | 50110.9163
T/CO2/Yr Saved | 141.849
Payback Period (Yrs) | 4.1907

## Sheet1
Design Feature | ‘Old’ Design Drying Cabinets | ‘New’ E3 Drying Cabinets
Insulation | None – Surfaces become very hot, unsafe to tough. High energy consumption. Higher HVAC costs. | Insulated – Both the control panel and chamber are insulated with natural materials. The doors of the larger units are also double glazed. Safe to touch and energy efficient. Lower HVAC costs
Temperature Setting | Dial – Numerical dial 1-10 made setting the exact temperature challenging, internal conditions could be too hot for contents (plasticware and tips). | Digital – Temperature can be precisely set to 0.1C to suit the contents being dried.
Temperature Control System | Simmerstat – Basic controller with poor temperature accuracy and variability at set point. | Microprocessor – Heating element is finely controlled, temperature is held accurately. Improves energy efficiency
Temperature Display | None – External monitoring is required to know what tm | L.E.D. – Temperature is clearly displayed.
Programmability | None - Units must be switched off manually or an external timer must be used. | 7 Day Timer – Unit will only be on when it’s required.